UT8805E Firmware upgrade guide

**1.** Copy the file " UT8805.UGD" to be upgraded to the USB flash drive (formatted as FAT32)

**NOTE:** BEFORE copying files to your USB. Verify USB Format is FAT32.

**If needed:** Format your USB to FAT32 format before use.

Newer USBs are usually NTFS file format and will not work.



1. Right click on USB drive. Select Format.
2. Select FAT32 from File system pull down menu.
3. Select Start

A screenshot of a computer


AI-generated content may be incorrect.

**2.** Insert the USB flash drive into the USB Host interface on the front panel of the multimeter

**3.** Press Shift→Dual→[Upgrade], press [Locate] and then press [Up] or [Down] to select the upgrade file, then press [Select]→[Done] to start the upgrade

**4.** After the upgrade is completed, restart the instrument and then remove the USB flash drive.



**5.** Restart the device and check the upgrade version information

Press the Shift key and then the Dual key, then select [System] → [About] to check whether the upgraded software and hardware version numbers are consistent with the target version.

If they are inconsistent, the upgrade is unsuccessful and you need to follow the above steps to upgrade again.



**6.** After checking, press [Done] to exit the system information interface.

**Notice:**

During the upgrade process, please do not pull out the USB flash drive, shut down or lose power!